## Dashboard
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL
Medical Equipment Management System — Summary Dashboard
INVENTORY STATISTICS
Total Equipment Records (WHO Definition) |  | 170
Functional |  | 153
Functional (New) |  | 0
Non-Functional |  | 12
To Be Disposed |  | 5
Transferred |  | 0
WORKBOOK CONTENTS
Equipment Inventory | 170 WHO-classified equipment items from 2018 E.C. inventory
Performance Verification | 13 PM/verification checklists — Oxygen Concentrators & Mechanical Ventilators (July 2016)
Work Orders | 13 corrective maintenance work orders
Preventive Maintenance | 2 planned PM task reports
Acceptance Testing | Structured acceptance test template based on Menelik II form
Training Records | 12 personnel training verification entries (Oct 2015)
Calibration | Calibration log — June 2017 E.C. report reference (details to be added)

## Equipment Inventory
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — Medical Equipment Inventory (2018 E.C.)
Source: 2018 E.C. First Round Inventory ¦ Filtered: WHO Essential Medical Equipment (electrical/electronic only; excludes mechanical scales, exam lights, passive tools) ¦ Total: 170
# | Inventory Number | Equipment Name (Standardized) | WHO Category | Original Name in Record | Department / Ward | Manufacturer / Brand | Model | Serial Number | Status
1 | AA/DMCH/104/4523/new/CB/MED/A/0001 | Patient Monitor/Screen | Patient Monitoring | SCREEN,Patient | 3rd floor |  |  |  | Functional
2 | AA/DMCH/104/4523/833/CB/MED/A/0001 | Oxygen Concentrator | Life Support & Respiratory | Oxygen concentrator | 3rd floor | yuye |  |  | Functional
3 | AA/DMCH/104/4523/DBP/CB/MED/A/0001 | Digital BP Monitor | Vital Signs & Diagnostic | BP cuff,DIGITAL | 3rd floor | ROSSMAX | 196025000024 | 196025000024 | Functional
4 | AA/DMCH/104/4523/DBP/CB/MED/A/0002 | Digital BP Monitor | Vital Signs & Diagnostic | BP cuff, DIGITAL | 3rd floor | RIESTER | 000897 | 000897 | Functional
5 | AA/DMCSH/104/4523/Cryo/CB/ABOR/0001 | Cryotherapy Machine | Physiotherapy & Rehabilitation | Cryotherapy Machine | Abortion(Family Planning) | mgc-200 | 17712 | 17712 | Functional
6 | AA/DMCSH/104/4523/210/CB/EM/0002 | Patient Monitor | Patient Monitoring | Monitor, Patient | Adult Emergency | Anyview A8 | A027E002473 | A027E002473 | Functional
7 | AA/DMCSH/104/4523/210/CB/EM/0001 | Patient Monitor | Patient Monitoring | Monitor, Patient | Adult Emergency | TM-900 | C721081867 | C721081867 | Non-Functional
8 | AA/DMCSH/104/4523/210/CB/EM/0005 | Patient Monitor | Patient Monitoring | Monitor, Patient | Adult Emergency | Anyview A8 | A027E002103 | A027E002103 | Functional
9 | AA/DMCSH/104/4523/775/CB/EM/0002 | Dry Oven | Laboratory Equipment | Dry Oven | Adult Emergency | YCO-010 | 618529 | 618529 | Functional
10 | AA/DMCSH/104/4523/799/CB/EM/0004 | Infusion Pump | Life Support & Respiratory | Infusion- Pump | Adult Emergency | HK-100 | 01111049 | 01111049 | Functional
11 | AA/DMCSH/104/4523/799/CB/EM/0005 | Infusion Pump | Life Support & Respiratory | Infusion- Pump | Adult Emergency | HK-100 | 01111247 | 01111247 | Functional
12 | AA/DMCSH/104/4523/799/CB/EM/0006 | Infusion Pump | Life Support & Respiratory | Infusion- Pump | Adult Emergency | HK-100 | 01111041 | 01111041 | Functional
13 | AA/DMCSH/104/4523/811/CB/ICU/0002 | Mechanical Ventilator, Mobile | Life Support & Respiratory | Mechanical Ventilator, Mobile | Adult Emergency | MV2000 | 08809268914133 | 08809268914133 | Functional
14 | AA/DMCSH/104/4523/811/CB/ICU/0001 | Mechanical Ventilator, Mobile | Life Support & Respiratory | Mechanical Ventilator, Mobile | Adult Emergency | SV300 | GB-26054191 | GB-26054191 | Functional
15 | AA/DMCSH/104/4523/833/CB/EM/0003 | Oxygen Concentrator | Life Support & Respiratory | Oxygen Concentrator | Adult Emergency | Yuyue |  |  | Functional
16 | AA/DMCSH/104/4523/833/CB/EM/0052 | Oxygen Concentrator | Life Support & Respiratory | Oxygen Concentrator | Adult Emergency | Newlife Intensity | CBB0119230246 | CBB0119230246 | Functional
17 | AA/DMCSH/104/4523/127/CB/EM/0004 | Suction Machine, Electrical | Life Support & Respiratory | Suction machine, Electrical | Adult Emergency | Gima | 1096 | 1096 | Functional
18 | AA/DMCSH/104/4523/889/CB/IMG/0006 | Ultrasound Machine | Diagnostic Imaging | Ultrasound | Adult Emergency | DC13 | CA1/91001149 | CA1/91001149 | Functional
19 | AA/DMCSH/104/4523/902/CB/EM/0001 | X-Ray View Box, Powered | Diagnostic Imaging | X-ray view box, powered | Adult Emergency |  |  |  | Functional
20 | AA/DMCSH/104/4523/50/CB/EM/0004 | Steam Sterilizer | Sterilization Equipment | Sterilizer, Steam | Adult Emergency | SA300H | 2010240011-002 | 2010240011-002 | Functional
21 | AA/DMCSH/104/4523/152/CB/EOPD/0007 | X-Ray Machine, Mobile Digital | Diagnostic Imaging | X-Ray, Mobile, Digital | Adult Emergency | Mobile Dart | 41BDDF944001 | 41BDDF944001 | Functional
22 | AA/DMCSH/104/4523/152/CB/EOPD/0010 | Suction Machine, Electrical | Life Support & Respiratory | Suction machine, Electrical | Adult Emergency | ASKIR C30 AP FS 2 | 69064 | 69064 | Non-Functional
23 | AA/DMCSH/104/4523/130/CB/EM/0001 | Defibrillator | Surgery & Theatre | Defibrilator | Adult Emergency |  | DGB5B085259 | DGB5B085259 | Functional
24 | AA/DMCSH/104/4523/776/CB/EM/0003 | ECG Machine | Vital Signs & Diagnostic | ECG Machine | Adult Emergency | MAC 600 | SS521360015PA | SS521360015PA | To Be Disposed
25 | AA/DMCSH/104/4523/152/CB/EM/0008 | Suction Machine, Electrical | Life Support & Respiratory | Suction machine, Electrical | Adult Emergency | 7A-23D | 2311110113 | 2311110113 | Functional
26 | AA/DMCSH/104/4523/130/CB/EM/0002 | Defibrillator | Surgery & Theatre | Defibrilator | Adult Emergency | binheart-03 | ENSB023254 | ENSB023254 | Functional
27 | AA/DMCSH/104/4523/762/CB/ARTLAB/0001 | Centrifuge | Laboratory Equipment | CENTRIFUGE | ART LAB | TD6C | 24661905039 | 24661905039 | Functional
28 | AA/DMCSH/104/4523/723/CB/ARTLAB/0001 | CD4 Analyzer | Laboratory Equipment | CD4 Machine | ART LAB | BDFACS PRESTO | M6631100000989 | M6631100000989 | Functional
29 | AA/DMCH/104/4523/751/EYEOPD/CB/0001 | Auto Refractometer | Ophthalmology | Autorefractometer | biometry room(eye opd) | AR/205--- | 141358 | 141358 | Functional
30 | AA/DMCH/104/4523/889/EYEOPD/CB/0001 | Ultrasound Machine | Diagnostic Imaging | Ultrasound | biometry room(eye opd) | quntel medical | 3628 | 3628 | Functional
31 | AA/DMCSH/104/4523/170/CB/LAB/0001 | Water Bath | Laboratory Equipment | Water Bath | Blood Bank | YCW010 | 608259 | 608259 | Functional
32 | AA/DMCH/104/4523/933/BB/CB/0001 | Platelet Agitator | Laboratory Equipment | Platelet agitator | blood bank | BPA-01 | 1480 | 1480 | Non-Functional
33 | AA/DMCH/104/4523/792/BB/CB/0001 | Plasma Freezer | Laboratory Equipment | plazma freezer | blood bank | Helmer | 2108025 | 2108025 | Functional
34 |  | Microscope, Light | Laboratory Equipment | Microscope | blood bank | CX-21 | 0633508 | 0633508 | Functional
35 |  | Centrifuge | Laboratory Equipment | CENTRIFUGE | blood bank | R-8CPLUS | ZGEN-10054 | ZGEN-10054 | Functional
36 | AA/DMCH/104/4523/933/BB/CB/0001 | Plasma/Platelet Agitator | Laboratory Equipment | Plazma agitator | blood bank | Helmer | 2074349 | 2074349 | Functional
37 | AA/DMCSH/104/4523/50/CB/CS/0001 | Autoclave, Steam | Sterilization Equipment | Autoclave | Central Steralization | SAKURA | 1C-37-B | 1C-37-B | Functional
38 | AA/DMCSH/104/4523/50/CB/CS/0002 | Autoclave, Steam | Sterilization Equipment | Autoclave | Central Steralization | CIS42104/04/1 | 23996 | 23996 | Functional
39 | AA/DMCSH/104/4523/50/CB/CS/0003 | Autoclave, Steam | Sterilization Equipment | Autoclave | Central Steralization | 5596-1V Tuttnauer | 20100024 | 20100024 | Functional
40 | AA/DMCSH/104/4523/DBP/CB/CNT/0001 | Digital BP Monitor | Vital Signs & Diagnostic | BP CUFF DIGITAL | Central Triage | RBP-100 | 6843 | 6843 | Functional
41 | AA/DMCSH/104/4523/1164/CB/COPH/0001 | Analytical Balance, Digital | Laboratory Equipment | Analytical balance, digital | Compound Pharmacy | BN112BB | 059810 | 059810 | Functional
42 | AA/DMCSH/104/4523/774/CB/COPH/0001 | Water Distiller | Laboratory Equipment | Distiller | Compound Pharmacy | WDZ-20 | 12001643 | 12001643 | To Be Disposed
43 | AA/DMCSH/104/4523/distil/CB/cop ph/0001 | Water Distiller | Laboratory Equipment | Distiller | Compound Pharmacy | GYZD-10 |  |  | Functional
44 | AA/DMCSH/104/4523/127/CB/LW/0001 | Suction Machine | Life Support & Respiratory | Suction Machine | Delivery room | 7A-23D | 23122200017 | 23122200017 | Functional
45 | AA/DMCH/104/4523/NEW/CB/LW/0001 | Infant Radiant Warmer | Neonatology | infant radiant warmer | Delivery room | BN-100 | 1130105268 | 1130105268 | Functional
46 | AA/DMCH/104/4523/NEW/CB/LW/0001 | Fetal Monitor | Patient Monitoring | fetal moniter | Delivery room |  |  |  | Functional
47 | AA/DMCH/104/4523/NEW/CB/LW/0002 | Fetal Monitor | Patient Monitoring | fetal moniter | Delivery room |  |  |  | Functional
48 | AA/DMCSH/104/4523/770/CB/DN/0001 | Dental Unit | Dental Equipment | Dental Unit | Dental Room | mikrodental E-line 100 | 1901-pc052 | 1901-pc052 | Functional
49 | AA/DMCSH/104/4523/770/CB/DN/0002 | Dental Unit | Dental Equipment | Dental Unit | Dental Room | mikrodental E-line 100 | 1901-pc045 | 1901-pc045 | Functional
50 | AA/DMCSH/104/4523/770/CB/DN/0003 | Dental Unit | Dental Equipment | Dental Unit | Dental Room | mikrodental E-line 100 | not tagged | not tagged | Functional
51 | AA/DMCH/104/4523/145/CB/DERMA/0001 | Pulse Oximeter | Patient Monitoring | Pulse oximeter | Dermathology | RAD-G | 600001683 | 600001683 | Functional
52 | AA/DMCSH/104/4523/DIA/CB/DIA/CENTER/0010 | Dialysis Machine | Renal Dialysis | Dialysis machine | dialysis center | surdial55 plus | 23CX30452 | 23CX30452 | Functional
53 | AA/DMCSH/104/4523/DIA/CB/DIA/CENTER/0011 | Dialysis Machine | Renal Dialysis | Dialysis machine | dialysis center | surdial55 plus | 23CX29048 | 23CX29048 | Functional
54 | AA/DMCSH/104/4523/DIA/CB/DIA/CENTER/0012 | Dialysis Machine | Renal Dialysis | Dialysis machine | dialysis center | surdial55 plus | 23CX29516 | 23CX29516 | Functional
55 | AA/DMCSH/104/4523/145/CB/DIA/CENTER/0001 | Pulse Oximeter | Patient Monitoring | Pulse oximeter | dialysis center | Rad-G | 6000054409 | 6000054409 | Functional
56 | AA/DMCSH/104/4523/127/CB/DIA/CENTER/0001 | Suction Machine | Life Support & Respiratory | Suction machine | Dialysis center | 7A-23D | 23122200062 | 23122200062 | Functional
57 | AA/DMCSH/104/4523/889/CB/DIA/CENTER/0001 | Ultrasound Machine | Diagnostic Imaging | Ultrasound | Dialysis center | Z 50 | HH6-17001361 | HH6-17001361 | Functional
58 | AA/DMCSH/104/4523/924/CB/ENT/0001 | ENT Chair (Powered) | ENT Equipment | ENT chair | ENT |  |  |  | Functional
59 | AA/DMCSH/104/4523/787/CB/ENT/0001 | ENT Chair (Powered) | ENT Equipment | ENT Chair | ENT | apex-2000 |  |  | Functional
60 | AA/DMCSH/104/4523/127/CB/ENT/0001 | Suction Machine | Life Support & Respiratory | Suction machine | ENT |  |  |  | Functional
61 | AA/DMCSH/104/4523/775/CB/EEM/0001 | Dry Oven | Laboratory Equipment | Dry Oven | EYE Emergency | E-34 | 84278 | 84278 | Functional
62 | AA/DMCSH/104/4523/779/CB/EEM/0001 | Electrosurgical Unit | Surgery & Theatre | Electrosurigcal Unit | EYE Emergency | diatrom 300 | 4000161956 | 4000161956 | Functional
63 | AA/DMCSH/104/4523/855/CB/EOPD/0001 | Slit Lamp | Ophthalmology | Slit Lamp | EYE Emergency |  | 3702066 | 3702066 | Functional
64 | AA/DMCSH/104/4523/876/CB/EEM/0005 | Tonometer | Ophthalmology | Tonometer | EYE Emergency | Keeler | 2414-110457 | 2414-110457 | Functional
65 | AA/DMCH/104/4523/855/EYE/EM/CB/0001 | Slit Lamp | Ophthalmology | Slit Lamp | EYE Emergency | SLM-3ER | 68792204 | 68792204 | Functional
66 | AA/DMCH/104/4523/NEW/EYE/EM/CB/0001 | Patient Monitor/Screen | Patient Monitoring | SCREEN,Patient | EYE Emergency |  |  |  | Functional
67 | AA/DMCSH/104/4523/751/CB/EOPD/0003 | Auto Refractometer | Ophthalmology | Auto Refractometer | Eye OPD | Visuref 100 | K85CA86D | K85CA86D | Functional
68 | AA/DMCSH/104/4523/751/CB/EOPD/0004 | Auto Refractometer | Ophthalmology | Auto Refractometer | Eye OPD | RMK - 150 | OP-RMK-150A05717 | OP-RMK-150A05717 | Functional
69 | AA/DMCSH/104/4523/751/CB/EOPD/0005 | Biometry Unit | Ophthalmology | Biometry | Eye OPD | IOLMASTER | 1000543 | 1000543 | Functional
70 | AA/DMCSH/104/4523/751/CB/EOPD/0006 | Fundus Camera | Diagnostic Imaging | Fundus Camera | Eye OPD | AAFCO1 | FC2311102 | FC2311102 | Functional
71 | AA/DMCSH/104/4523/920/CB/EOPD/0001 | Green Laser (Ophthalmic) | Ophthalmology | Green Laser | Eye OPD | AMO9HSMART | ILS23081048 | ILS23081048 | Functional
72 | AA/DMCSH/104/4523/855/CB/EOPD/0001 | Slit Lamp | Ophthalmology | Slit Lamp | Eye OPD | 8127004 | 1910-03212 | 1910-03212 | Functional
73 | AA/DMCSH/104/4523/876/CB/EOPD/0001 | Tonometer | Ophthalmology | Tonometer | Eye OPD | Visuplan 500 | 9703102463 | 9703102463 | Functional
74 | AA/DMCSH/104/4523/876/CB/EOPD/0002 | Tonometer | Ophthalmology | Tonometer | Eye OPD | Visuplan 500 | 9703102470 | 9703102470 | Functional
75 | AA/DMCH/104/4523/912/GL/CB/0002 | Laser Machine (Ophthalmic) | Ophthalmology | Lazer machine | EYE OPD | LSLYA 9111 | 1208846 | 1208846 | Functional
76 | AA/DMCH/104/4523/855/EYEOPD/CB/0001 | Corneal Topography | Ophthalmology | Corneal Topography | EYE OPD | 9000 | 9000-103776 | 9000-103776 | Functional
77 | AA/DMCH/104/4523/855/EYEOPD/CB/0002 | Indirect Ophthalmoscope | Ophthalmology | indarect opthalomoscope | EYE OPD | 7-43 | 74301 | 74301 | Non-Functional
78 | AA/DMCSH/104/4523/743/CB/EOR/0003 | Anesthesia Machine | Life Support & Respiratory | Anesthesia Machine | Eye OR | Pro-45 |  |  | Functional
79 | AA/DMCSH/104/4523/743/CB/EOR/0004 | Anesthesia Machine | Life Support & Respiratory | Anesthesia Machine | Eye OR | Aeon7200 | Aeon7200€yxz5044 | Aeon7200€yxz5044 | Functional
80 | AA/DMCSH/104/4523/766/CB/EOR/0001 | Cryotherapy Machine | Physiotherapy & Rehabilitation | Cyromatic | Eye OR | MKII | K0909825 | K0909825 | Functional
81 | AA/DMCSH/104/4523/818/CB/EOR/0004 | Microscope, Operating Ophthalmic | Laboratory Equipment | Microscope, Operating, Ophthalmic | Eye OR | AAOM230/250 Brilliant | 1462500350 | 1462500350 | Functional
82 | AA/DMCSH/104/4523/818/CB/EOR/0008 | Microscope, Operating Ophthalmic | Laboratory Equipment | Microscope, Operating, Ophthalmic | Eye OR | Opmi Lumera 300 | 6137102904 | 6137102904 | Functional
83 | AA/DMCSH/104/4523/818/CB/EOR/0003 | Microscope, Operating Ophthalmic | Laboratory Equipment | Microscope, Operating, Ophthalmic | Eye OR | Opmi Lumera 300 | 6137102905 | 6137102905 | Functional
84 | AA/DMCSH/104/4523/870/CB/EOR/0006 | Operating Table, Electro-Hydraulic | Surgery & Theatre | Table, OR, Electro hydraulic | Eye OR | OP830 | OP830YUYY | OP830YUYY | Functional
85 | AA/DMCSH/104/4523/743/CB/EOR/0011 | Anesthesia Machine | Life Support & Respiratory | Anesthesia Machine | Eye OR | 7200 | Gloryplus(s)yuyz094 | Gloryplus(s)yuyz094 | Non-Functional
86 | AA/DMCH/104/4523/818/CB/EYEOR/0005 | Microscope, Operating | Laboratory Equipment | Microscope,operating | eye OR | S8 | 6627502771 | 6627502771 | Functional
87 | AA/DMCH/104/4523/818/CB/EYEOR/0006 | Microscope, Operating | Laboratory Equipment | Microscope,operating | eye OR | S8 | 6629184942 | 6629184942 | Functional
88 | AA/DMCH/104/4523/PHECO/CB/EYEOR/0006 | Phaco Machine | Surgery & Theatre | PHECO MACHINE | eye OR | VC821100 | 73700084 | 73700084 | Functional
89 | AA/DMCH/104/4523/IRLAZ/CB/EYEOR/0001 | Appasamy Infrared Laser | Ophthalmology | Appasamy infrared laser | eye OR | smart-810NM | IRL2312008 | IRL2312008 | Functional
90 | AA/DMCH/104/4523/PHECO/CB/EYEOR/0002 | PPV + Phaco Machine | Surgery & Theatre | PPV + PHECO MACHINE | eye OR | GALAXY TURBO ORBIT | 6304230093 | 6304230093 | Functional
91 | AA/DMCH/104/4523/775/CB/EYEOR/0007 | Hot Air Sterilizer | Sterilization Equipment | STERALIZER,HOT AIR | eye OR | YCO-010 | 618527 | 618527 | Functional
92 | AA/DMCH/104/4523/775/CB/EYEOR/0003 | Hot Air Sterilizer | Sterilization Equipment | Steralizer,hot air | eye OR | YCO-010 | 618312 | 618312 | Functional
93 | AA/DMCH/104/4523/775/CB/EYEOR/0005 | Hot Air Sterilizer | Sterilization Equipment | Steralizer,hot air | eye OR | DE-00T3 | 0112 | 0112 | Functional
94 | AA/DMCH/104/4523/818/CB/EYEOR/0001 | Microscope, Operating | Laboratory Equipment | Microscope,operating | eye OR | S8 | 6629184951 | 6629184951 | Functional
95 | AA/DMCH/104/4523/775/CB/EYEOR/0006 | Dry Oven | Laboratory Equipment | Dry oven | eye OR |  |  |  | Functional
96 | AA/DMCH/104/4523/PHECO/CB/EYEOR/0005 | PPV + Phaco Machine | Surgery & Theatre | ppv+Pheco machine | eye OR | GALAXY TURBO ORBITAL | 6304230094 | 6304230094 | Functional
97 | AA/DMCH/104/4523/PHECO/CB/EYEOR/00010 | Phaco Machine | Surgery & Theatre | PHECO MACHINE | eye OR | PULSAR | 65AQ | 65AQ | Functional
98 | AA/DMCH/104/4523/PHECO/CB/EYEOR/0005 | PPV + Phaco Machine | Surgery & Theatre | ppv+Pheco machine | eye OR | V500 |  |  | Functional
99 | AA/DMCH/104/4523/145/CB/EYEOR/0008 | Pulse Oximeter | Patient Monitoring | Pulse oximeter | eye OR | (70-1065) | GIB008051 | GIB008051 | Functional
100 | AA/DMCH/104/4523/50/CB/EYEOR/0009 | Steam Sterilizer | Sterilization Equipment | Steam sterialzer | eye OR | CISA |  |  | Non-Functional
101 | AA/DMCH/104/4523/779/CB/EYEOR/0043 | Electrosurgical Unit | Surgery & Theatre | Electrosurigcal Unit | eye OR | avanta | L223B02AN | L223B02AN | Functional
102 | AA/DMCH/104/4523/818/CB/EYEOR/0001 | Bipolar Coagulator | Surgery & Theatre | Bipolar coagulator | eye OR | ACC001 | 2710150017 | 2710150017 | Functional
103 | AA/DMCH/104/4523/818/CB/EYEOR/0044 | Electrosurgical Unit | Surgery & Theatre | Electrosurigcal Unit | eye OR | ACD01 | E14G-05 | E14G-05 | Functional
104 | AA/DMCH/104/4523/818/CB/EYEOR/0005 | Bipolar Coagulator | Surgery & Theatre | Bipolar coagulator | eye OR | ACD01 | EAE-17 | EAE-17 | Functional
105 | AA/DMCH/104/4523/NEW/CB/EYEOR/0001 | Vitrectomy Machine | Surgery & Theatre | VITROCTOMY | eye OR | G-23004 | VIT-1748444 | VIT-1748444 | Functional
106 | AA/DMCH/104/4523/NEW/CB/EYEOR/0002 | Vitrectomy Machine | Surgery & Theatre | Vitroctomy | eye OR | G-23004 | VIT-0523019 | VIT-0523019 | Functional
107 | AA/DMCSH/104/4523/50/CB/CX-CAS/0001 | Steam Sterilizer | Sterilization Equipment | Steam steralizer | family planning | AS-300H | 307388 | 307388 | Functional
108 | AA/DMCH/104/4523/912/GL/CB/0001 | Laser Machine (Ophthalmic) | Ophthalmology | Lasermachine | Gelacoma | SL990-3X | 13070306 | 13070306 | Functional
109 | AA/DMCH/104/4523/912/GL/CB/0010 | YAG Laser Machine | Ophthalmology | yaglasermachine | Gelacoma | ESL-1 | 4290AR | 4290AR | Functional
110 | AA/DMCSH/104/4523/125/CB/GYOR/0007 | Infant Radiant Warmer | Neonatology | Radaint Warmer | Gyni OR | BN-100 | 1130105118 | 1130105118 | Functional
111 | AA/DMCSH/104/4523/125/CB/GYOR/0012 | Infant Radiant Warmer | Neonatology | Radaint Warmer | Gyni OR | BN-100 | 1130105140 | 1130105140 | Functional
112 | AA/DMCSH/104/4523/170/CB/LAB/0003 | Water Bath | Laboratory Equipment | Water Bath | Gyni OR |  |  |  | Functional
113 | AA/DMCSH/104/4523/870/CB/GYOR/0007 | Operating Table, Electric | Surgery & Theatre | Table, OR, Electric | Gyni OR | FE 300 | 2016-10 | 2016-10 | Functional
114 | AA/DMCSH/104/4523/870/CB/GYOR/0001 | Operating Table, Electro-Hydraulic | Surgery & Theatre | Table, OR, Electro hydraulic | Gyni OR | OP330 |  |  | Functional
115 | AA/DMCSH/104/4523/130/CB/ICU/0002 | Defibrillator | Surgery & Theatre | Defibrilator | ICU | Benehart D3 | EL-53019652 | EL-53019652 | Functional
116 | AA/DMCSH/104/4523/811/CB/ICU/0002 | Mechanical Ventilator, Mobile | Life Support & Respiratory | Mechanical Ventilator, Mobile | ICU | SV-300 |  |  | Functional
117 | AA/DMCSH/104/4523/735/CB/ICU/0001 | Blood Gas Analyzer | Laboratory Equipment | Analyzer, Blood Gas | ICU | ABL800flex | YZB/DEN6769-2013 | YZB/DEN6769-2013 | To Be Disposed
118 | AA/DMCH/104/4523/811//CB/ICU/0008 | Mechanical Ventilator | Life Support & Respiratory | Mechanical ventilator | ICU | V200 | VS0003162 | VS0003162 | Non-Functional
119 | AA/DMCH/104/4523/811//CB/ICU/0009 | Mechanical Ventilator | Life Support & Respiratory | Mechanical ventilator | ICU | V200 | VS0003284 | VS0003284 | Non-Functional
120 | AA/DMCH/104/4523/811//CB/ICU/0010 | Mechanical Ventilator | Life Support & Respiratory | Mechanical ventilator | ICU | V200 | VS0003223 | VS0003223 | Non-Functional
121 | AA/DMCSH/104/4523/1160/CB/IM/0001 | Automatic Injector | Surgery & Theatre | Automatic Injector | Imaging | Dual Shot Alpha | QCBS0011G | QCBS0011G | Functional
122 | AA/DMCSH/104/4523/910/CB/IMG/0001 | X-Ray Machine, Radiography Digital | Diagnostic Imaging | X-Ray, Radiography, Digital | Imaging | New Orental 1000 | y13/284/15/9 | y13/284/15/9 | Functional
123 | AA/DMCSH/104/4523/767/CB/IMG/0001 | CT Scanner | Diagnostic Imaging | CT Scan | Imaging | 5122080-12 | 362870HMB | 362870HMB | Non-Functional
124 | AA/DMCSH/104/4523/152/CB/EOPD/0008 | X-Ray Machine, Radiography | Diagnostic Imaging | X-Ray, Radiography | Imaging | 1/2P13DK | 83697 | 83697 | Functional
125 | AA/DMCSH/104/4523/152/CB/EOPD/0009 | X-Ray Machine w/ Fluoroscopy | Diagnostic Imaging | X-Ray, Radiography, Analog with fluoroscopy | Imaging | New Orental 1001 | y13/284/15/8 | y13/284/15/8 | Non-Functional
126 | AA/DMCH/104/4523/816/CB/LAB/0001 | Microscope, Light | Laboratory Equipment | Microscope | LAB,Emergency | Primo Star | 3144005621 | 3144005621 | Functional
127 | AA/DMCH/104/4523/762/CB/LAB/0001 | Centrifuge, Analytical | Laboratory Equipment | centrifuge,analytical | LAB,Emergency | CATNG0151 | 169077 | 169077 | Functional
128 | AA/DMCSH/104/4523/163/CB/LAB/0001 | Hematology Analyzer | Laboratory Equipment | Analyzer, Hematology | Laboratory | BC-5150 | 25006175 | 25006175 | Functional
129 | NEW | Biological Safety Cabinet | Laboratory Equipment | Biological Safety | Laboratory | BSC-1000IIAR |  |  | Functional
130 | NEW | Magnetic Stirrer | Laboratory Equipment | Biometric Strvie magnetic stirrer | Laboratory |  | MS00604 | MS00604 | Functional
131 | AA/DMCSH/104/4523/762/CB/LAB/0002 | Centrifuge | Laboratory Equipment | Centrifuge | Laboratory | R-8C Plus | ZGEN-10056 | ZGEN-10056 | Functional
132 | NEW | Densitometer | Laboratory Equipment | Densometer | Laboratory | VEN 1 | 00010223010068 | 00010223010068 | Functional
133 | AA/DMCSH/104/4523/797/CB/LAB/0001 | Incubator, Infant | Neonatology | Incubator | Laboratory | IN-010-helmer | 2074349 | 2074349 | Functional
134 | AA/DMCSH/104/4523/816/CB/LAB/0001 | Microscope, Light | Laboratory Equipment | Microscope, Light | Laboratory | Primo Star | 3149016163 | 3149016163 | Functional
135 | NEW | Plasma Thawing System | Laboratory Equipment | Plasma Thawing system | Laboratory | HELMER | 2106856 | 2106856 | Functional
136 | AA/DMCSH/104/4523/762/CB/EOPD/0007 | Centrifuge, Urine | Laboratory Equipment | Urine electrical centrifuge | Laboratory | AVI-526 | 92026.092718 | 92026.092718 | Functional
137 | AA/DMCSH/104/4523/170/CB/LAB/0002 | Water Bath | Laboratory Equipment | Water Bath | Laboratory | YCO-010 | 608259 | 608259 | Functional
138 | AA/DMCSH/104/4523/163/CB/LAB/0002 | Hematology Analyzer | Laboratory Equipment | Analyzer, Hematology | Laboratory | BC-S150 | SR-25006176 | SR-25006176 | Functional
139 | AA/DMCSH/104/4523/774/CB/LAB/0002 | Water Distiller | Laboratory Equipment | Distiller | Laboratory | WDZ20 | 12001601 | 12001601 | Functional
140 | AA/DMCSH/104/4523/794/CB/LAB/0001 | GeneXpert Machine | Laboratory Equipment | Gene Expert | Laboratory | Cepheid | 809807 | 809807 | Functional
141 | AA/DMCSH/104/4523/1162/CB/LAB/0002 | Shaker | Laboratory Equipment | Shaker | Laboratory | VRN-200 | 509890 | 509890 | Functional
142 | AA/DMCSH/104/4523/83/CB/LAB/0003 | Clinical Chemistry Analyzer | Laboratory Equipment | Analzer, Clinical, Chemistry | Laboratory | BS-200 |  |  | Non-Functional
143 | AA/DMCH/104/4523/NEW/LAB/CB/0001 | Coagulation Analyzer | Laboratory Equipment | coagulation analyzer | laboratory | XL-3600C |  |  | Functional
144 | AA/DMCH/104/4523/83/LAB/CB/0001 | Clinical Chemistry Analyzer | Laboratory Equipment | Chemistry analyzer | laboratory | 240 | 010102000702200306 | 010102000702200306 | Functional
145 | AA/DMCH/104/4523/BBD/CB/LAB/0001 | Analytical Balance, Digital | Laboratory Equipment | Bean balance,digital | laboratory | AL204-IC | 1128253384 | 1128253384 | Functional
146 | AA/DMCSH/104/4523/870/CB/MJOR/0003 | Operating Table, Electric | Surgery & Theatre | Table, OR, Electric | Major OR | Fazzini |  |  | Functional
147 | AA/DMCSH/104/4523/870/CB/MJOR/0005 | Operating Table, Electro-Hydraulic | Surgery & Theatre | Table, OR, Electro hydraulic | Major OR | Op830 |  |  | Functional
148 | AA/DMCSH/104/4523/784/CB/MJOR/0001 | Endoscopy System | Surgery & Theatre | Endoscopy | Major OR | LMD276NB | 300924 | 300924 | Functional
149 | AA/DMCSH/104/4523/LAPE/CB/MJOR/0001 | Laparoscopy System | Surgery & Theatre | laparascopy | Major OR | LMD-2765NB | 3000924 | 3000924 | Functional
150 | AA/DMCSH/104/4523/LAPE/CB/MJOR/0002 | Laparoscopy System | Surgery & Theatre | laparascopy | Major OR | V24.N1418 | C18-007645 | C18-007645 | Functional
151 | AA/DMCSH/104/4523/LAPE/CB/MJOR/0003 | Laparoscopy System | Surgery & Theatre | laparascopy | Major OR | stryeker | CVE-2111005 | CVE-2111005 | Functional
152 | AA/DMCSH/104/4523/1167/CB/MED/REF/0001 | EEG Machine | Patient Monitoring | EEG | Medical Refral | EEG-1200K | 4734 | 4734 | Functional
153 | AA/DMCH/104/4523/1162/MBIO/CB/0001 | Orbital Shaker | Laboratory Equipment | Orbital Shaker | microbiology | KS-20LBD |  |  | Functional
154 | AA/DMCH/104/4523/797/MBIO/CB/0001 | Incubator, Infant | Neonatology | electrical incubator | microbiology |  | 5-1481 | 5-1481 | Functional
155 | AA/DMCH/104/4523/797/MBIO/CB/0002 | Incubator, Lab | Laboratory Equipment | LAB Incubator | microbiology | IN-010 | 1807712 | 1807712 | Functional
156 | AA/DMCH/104/4523/797/MBIO/CB/0003 | Incubator, Lab | Laboratory Equipment | LAB Incubator | microbiology | IN-010 | 618571 | 618571 | Functional
157 | AA/DMCH/104/4523/797/MBIO/CB/0004 | Incubator, Infant | Neonatology | Incubator | microbiology | ARISHKAR | 410 | 410 | Functional
158 | AA/DMCH/104/4523/MST/CB/MBIO/0001 | Magnetic Stirrer/Stainer | Laboratory Equipment | Magnetic stainer | microbiology | uayledental | stno-m50064 | stno-m50064 | Functional
159 | AA/DMCH/104/4523/794/CB/MBIO/0001 | GeneXpert Machine | Laboratory Equipment | genexpert | microbiology | XPERTMTB | 809807 | 809807 | Functional
160 | AA/DMCSH/104/4523/144/CB/NICU/0005 | Phototherapy Unit | Neonatology | Phototherapy | NICU | BILITRON | NAK/36285 | NAK/36285 | Functional
161 | AA/DMCSH/104/4523/144/CB/NICU/0003 | Phototherapy Unit | Neonatology | Phototherapy | NICU | XHZ-90L | B46BZZH070006 | B46BZZH070006 | Functional
162 | AA/DMCSH/104/4523/144/CB/NICU/0008 | Phototherapy Unit | Neonatology | Phototherapy | NICU | XHZ-90L | A46BB0041004 | A46BB0041004 | Functional
163 | AA/DMCSH/104/4523/210/CB/NICU/0050 | Vital Signs Monitor | Patient Monitoring | Monitor, Vital Sign | NICU | V6 | V026E005517 | V026E005517 | To Be Disposed
164 | AA/DMCSH/104/4523/210/CB/NICU/0052 | Vital Signs Monitor | Patient Monitoring | Monitor, Vital Sign | NICU | V6 | V026E005504 | V026E005504 | To Be Disposed
165 | AA/DMCSH/104/4523/902/CB/ORO/0002 | X-Ray View Box, Powered | Diagnostic Imaging | X-ray view box, powered | Orthopedic out patient |  |  |  | Functional
166 | AA/DMCSH/104/4523/902/CB/ORO/0003 | X-Ray View Box, Powered | Diagnostic Imaging | X-ray view box, powered | Orthopedic out patient |  |  |  | Functional
167 | AA/DMCH/104/4523/NEW/PAT/CB/0001 | Rotary Microtome | Laboratory Equipment | Rotery microtomy | pathology | HS2205 | CJHS1413-79HS2205 | CJHS1413-79HS2205 | Functional
168 | AA/DMCH/104/4523/NEW/PAT/CB/0002 | Rotary Microtome | Laboratory Equipment | Rotery microtomy | pathology | HS2205 | CJHS1414-50HS2205 | CJHS1414-50HS2205 | Functional
169 | AA/DMCH/104/4523/AST/PAT/CB/0001 | Auto Stainer | Laboratory Equipment | auto stainer | pathology |  |  |  | Functional
170 | AA/DMCSH/104/4523/nebu/CB/PDE/0001 | Nebulizer | Life Support & Respiratory | nebulizer | Pediatric Emergency | w003 | 20180225/10 | 20180225/10 | Functional

## Performance Verification
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — Performance Verification Checklists
Source: Physical paper records photographed 25 May 2026 ¦ Covers Preventive Maintenance sessions July 2016
# | Equipment Name | Manufacturer | Model | Serial Number | Maintenance Site | Date | Work Order Type | Duration | Overall Result | Technicians | Test Equipment Used | Remarks / Performance Summary
1 | Oxygen Concentrator | — | — | — | Operation Room | 26/07/2016 | Preventive Maintenance | 50 min | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | Battery alarm, power failure test, battery voltage satisfactory. Cycle pressure readings unsatisfactory (model-specific range).
2 | Oxygen Concentrator | Shenyang Canta Medical Tech | V8-WN-NS | 2448141050620 | Emergency Room | 26/07/2016 | Preventive Maintenance | 50 min | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | O2 purity satisfactory. Battery alarm & power failure alarm unsatisfactory. Cycle pressure readings unsatisfactory.
3 | Oxygen Concentrator | Sheng ang canta Medical Tech Co.Ltd | V8-WN-NS | 2448705051220 | Emergency | 26/07/2016 | Preventive Maintenance | 50 min | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | O2 purity & battery voltage satisfactory. Battery alarm, power failure alarm, cycle pressure unsatisfactory.
4 | Oxygen Concentrator | Sheng ang canta Medical Tech Co.Ltd | V8-WN-NS | 2448664051220 | Emergency | 26/07/2016 | Preventive Maintenance | 50 min | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | O2 purity & battery voltage satisfactory. Cycle pressure unsatisfactory.
5 | Oxygen Concentrator | Longfian SciTech | JAZ-5BW | N2J5S12324D | Emergency | 26/07/2016 | Preventive Maintenance | 50 min | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | Battery test load alarm pass. Cycle pressure readings unsatisfactory.
6 | Oxygen Concentrator | Yuyue | — | — | Medical Ward | 26/07/2016 | Preventive Maintenance | 50 min | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | O2 purity unsatisfactory on analyzer. Battery alarm & power failure alarm, battery voltage satisfactory.
7 | Oxygen Concentrator | Yuyue | — | — | Medical Ward | 26/07/2016 | Preventive Maintenance | 1 hour | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | O2 purity, battery alarm, power failure alarm, battery voltage satisfactory. Cycle pressure readings unsatisfactory.
8 | Oxygen Concentrator | Yuyue | — | — | Medical Ward | 25/07/2016 | Preventive Maintenance | 1 hour | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Oxygen analyzer (Maxtec AF17199058); Digital pressure meter (DMG-2G SN:3130002) | 5 of 8 units satisfactory for all parameters; 2 units unsatisfactory for O2 purity & cycle pressure on analyzer; 1 unit unsatisfactory for O2 purity & cycle pressure.
9 | Mechanical Ventilator | Mekics | V2000 | MV2000TA2J5763 | ICU | 25/07/2016 | Preventive Maintenance | 1 hour | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Gas flow analyzer (FLUKE SN:4179049) | O2 alarm, air failure alarm pass. Breathing circuit leak — N/A (anesthesia unit). O2 calibration pass. Tidal volume, breath rate, O2 concentration pass. Pressure readings unsatisfactory; parameters measured 11 times — average unacceptable per ISO 80601-2-12-2011(E).
10 | Mechanical Ventilator | Mekics | V2000 | MV2000UA1D1620 | ICU | 25/07/2016 | Preventive Maintenance | 1 hour | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Gas flow analyzer (FLUKE SN:4179049) | Similar results to above. FIO2, tidal volume, breath rate satisfactory. Pressure, frequency readings unsatisfactory.
11 | Mechanical Ventilator | Philips | V200 | VS0003162 | ICU | 25/07/2016 | Preventive Maintenance | 1 hour | Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Gas flow analyzer (FLUKE SN:4179049) | All parameters satisfactory per ISO 80601-2-12-2011(E).
12 | Mechanical Ventilator / Anesthesia Unit | Mindray | WATO EX_35 | KG-8C003261 | Operation Room | 25/07/2016 | Preventive Maintenance | 1 hour | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Gas flow analyzer (FLUKE SN:4179049) | O2 alarm, air failure alarm pass. O2 calibration N/A (anesthesia unit). Tidal volume 360ml (limit 270–330) — FAIL. PEEP pass.
13 | Mechanical Ventilator / Anesthesia Unit | Shenzen Comen | AX-400 | X42021219004B | Operation Room | 25/07/2016 | Preventive Maintenance | 1 hour | Partial Pass | Tebeje Beyene, Atnafu Gojo, Getie, Chernet, Mesfin | Gas flow analyzer (FLUKE SN:4179049) | O2 alarm pass; air failure alarm pass. Breathing circuit leak test pass. O2 calibration N/A. Tidal volume, breath rate, O2 concentration generally satisfactory. Some pressure deviations noted.

## Work Orders
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — Work Orders (Corrective Maintenance)
Source: Physical work order binder photographed 25 May 2026 ¦ Biomedical Engineering Case Team
WO# | Equipment Name | Model | Serial Number | Department | Location / Room | Reported Time | Issue Description | Critical Situation? | Usability Impact | Corrective Action Taken | Post-Repair Result | Completed By | Completion Date | Approved By | Approval Date
WO-001 | OR Light | JEG-130 | 17101133 | Major OR | OR | 05:00 | Device not powering on — fault detected | Yes | Unable to use | Check power supply connection; Self-test/functionality test | — | — | — | Hargwegn | —
WO-002 | Suction Machine | MaxSpeed 62 | 221780ASP | Major OR | Major OR | 09:00 | Loose connection of suction tube | Yes | Unable to use | (g) Tightening loose connection; (h) Replace/repair broken wiring | Pass | Mekdes & Birhanu | 16/06/2018 | Jinwork Alegne | 16/06/2018
WO-003 | Slit Lamp | TS-LED | 2002-03427 | T22 | Interior Segment | 06:00 | Potentiometer not working — unable to change intensity | Yes | Unable to use | (b) Adjust mechanical part; (d) Restore functionality (button/touch screen) | Pass | Mekdes & Birhanu | 18/06/2018 | Esegenea Tlg | 18/06/2018
WO-004 | Patient Bed | MOTEK | — | MCH Ward | MCH Ward 07 | 10:50 | Installation issue | — | — | Installation resolved | — | Abel Getachew, Gereusalem Mondero, Hosat Seid | 17/06/2018 | Ewebet M | 17/06/2018
WO-005 | Patient Bed | MOTEK | 0242034162 | MCH Ward | MCH Ward 07 | 10:50 | Installation issue | — | — | Installation resolved | — | Abel Getachew, Gereusalem Mondero, Hosat Seid | 17/06/2018 | Ewebet M | 17/06/2018
WO-006 | Bipolar Coagulator / Autocautery | AUTOCAUTERY/ACP01 | E17E17 | Eye OR | Room 3 | 1 week | Board (IC) is not working | No | Delay in treatment | (m) Other: Chase/replace another board | Pass | Eskandir | 07/05/2019 | Arkeis | 07/05/2019
WO-007 | Suction Machine | 74-23D | — | Midwifery OR | Room 2 | 03:15 | Motor stacked (jammed) | — | Unable to use | (f) Lubrication; (m) Other: Unscrew the motor and cleaned | Yes — Functional | Birhanu & Eetseder | 22/12/2017 | Birhan | —
WO-008 | Suction Machine | S2-6400A | 10/6029 | Eye OR | OR | — | Motor stacked; check for leakage | No | Unable to use; Reduced functionality | (j) Leak repair: Yes | Pass | Birhanu Sane | 05/02/2018 | Hyenas | 05/02/2018
WO-009 | Suction Machine | — | — | Eye OR | Room 1 | 03:00 | Unable to open bottle cover | Yes | Unable to use; Delay in treatment | (m) Other: Tube connection solved | Pass | Fiyew Gor | 20/01/2018 | Eshnader | 20/01/2018
WO-010 | Suction Machine | F-600 | 1820-22 | Gaxmi Ward | Room 01 | 15 mins | Device breakdown reported | — | — | — | — | — | 25/03/2018 | Luisma Abore | 25/03/2018
WO-011 | Suction Machine | 7A-23D | 2Mdo28-0002 | Major OR | Room 3 | 05:30 | Reduced functionality (issue blank) | — | Reduced functionality | — | — | — | — | — | —
WO-012 | Infant Radiant Warmer | BM-100 | 1130105268 | Delivery | — | 04:00 | Device issue reported — finding blank | — | Unable to use; Delay in treatment | Corrective action in progress | — | Daniel | — | — | —
WO-013 | Infusion Pump | — | — | — | — | — | Fusion [pump] breakdown | — | — | — | — | Kilsman Aborre | 30/04/2018 | Mekomnon | 30/04/2018

## Preventive Maintenance
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — Planned Preventive Maintenance Task Reports
Source: PM binder photographed 25 May 2026 ¦ Menelik II CSH Biomedical Engineering Case Team
PM# | Equipment Name | Model | Serial Number | Department | Inventory Number | PM Date | Tasks Performed | Frequency | Next PM Date | Completed By | Approved By | Approval Date
PM-001 | Patient Monitor | BPM-1010 | — | Medical Wards | — | 24/07/2017 | 1. Cleaned with gauze & alcohol 2. Checked parameters (setup/calibration) 3. Checked power source | Semi-Annual | 24/02/2018 | Daniel Ayele, Birhaniyis | Amenise Atef | 24/07/2017
PM-002 | Dialysis Monitor / Medical Monitor | Dilux | — | Medical Ward | — | 22/02/2019 | 1. Checked power source 2. Covered area and calibration check 3. Cleaned with saline | Annual | 22/02/2020 | Abe Abel Ayele, Birhan Severia | Lens Severia | 22/02/2019

## Acceptance Testing
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — Acceptance Test Log
Source: Acceptance Test Logsheet (blank template captured 05/09/2018) ¦ To be completed per equipment receipt
# | Equipment Name | Date Received | Supplied By | Company Contact | Rep Present? | Damage on Delivery? | Equipment Complete? | Spare Parts? | Accessories? | User Manual? | Operator Manual? | Assembly By | All Parts Available? | Parts Fit Together? | Functionality Test Passed? | Equipment Accepted? | Training Provided to Users? | Training Provided to Tech Staff? | Remarks
AT-001 | [Equipment Name] | [Date] | [Company Name] | [Contact] | Yes/No | Yes/No | Yes/No | Yes/No (list) | Yes/No (list) | Yes/No | Yes/No | [Company Name] | Yes/No | Yes/No | Yes/No | Yes/No | Yes/No | Yes/No | [Notes]
ℹ Note: This sheet is a structured template based on the Menelik II Acceptance Test Logsheet (dated 05/09/2018 E.C.). Add one row per equipment unit received. Fields marked Yes/No require confirmation. Acceptance must be certified by the Biomedical Engineering Directorate.

## Training Records
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — User & Biomedical Personnel Training Records
Source: Training Verification Form (photographed 25 May 2026) ¦ Biomedical Engineering Case Team
# | Department / Ward | Medical Device | Model | Manufacturer / Supplier | Trainee Name | Position | Trained By | Date of Training | Assessment / Review Date | Comments
1 | Adult Emergency | Oxygen Concentrator | — | — | Gimenachis Fepadu | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
2 | Adult Emergency | Oxygen Concentrator | — | — | Abrham Tipha | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
3 | Adult Emergency | Oxygen Concentrator | — | — | Daniel M. | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
4 | Adult Emergency | Oxygen Concentrator | — | — | Samuel Yolele | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
5 | Adult Emergency | Oxygen Concentrator | — | — | Tigrr Elrom | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
6 | Adult Emergency | Patient Monitor | — | — | Gimenachis Fepadu | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
7 | Adult Emergency | Patient Monitor | — | — | Abrham Tipha | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
8 | Adult Emergency | Patient Monitor | — | — | Daniel M. | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
9 | Adult Emergency | Patient Monitor | — | — | Samuel Yolele | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
10 | Adult Emergency | Suction Machine | — | — | Gimenachis Fepadu | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
11 | Adult Emergency | Suction Machine | — | — | Abrham Tipha | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —
12 | Adult Emergency | Suction Machine | — | — | Daniel M. | Nurse / Clinician | Mekaoned Robens | 02/10/2015 | — | —

## Calibration
MENELIK II COMPREHENSIVE SPECIALIZED HOSPITAL — Calibration Records
Source: 'Calibration Performance and Results for Medical Equipment' report (June 2017 E.C.) ¦ Photographed 25 May 2026
Cal# | Equipment Name | Model | Serial Number | Department | Inventory Number | Calibration Date | Calibration Type | Calibrating Body / Agency | Result | Calibration Certificate # | Next Calibration Due | Remarks
CAL-001 | [Equipment Name — from June 2017 E.C. report] | — | — | — | — | June 2017 (E.C.) | Performance Calibration | Menelik II CSH Biomedical Engineering | [See source document] | — | [12 months from calibration date] | Source: Physical calibration report cover page photographed May 2026. Full detail from internal report.
ℹ Note: The June 2017 E.C. Calibration Performance and Results report was identified from the photographed documents. The cover page was captured but internal pages were not individually photographed. Add calibration records here as data becomes available from that report and future calibration activities.